INSTALL SHEET  (STAND ALONE)		 		STATION Name:_____________

Local Date/Time:_____________________ 	GMT Date/Time:______________________________

Field Team: ____________________________________________________________________

Equipment

Sensor S/N:		_______________	Sensor Type: 	________________

DAS S/N:		_______________

Clock S/N:		_______________

Flash Disk 1 S/N:	_______________	Size:		________________

Flash Disk 2 S/N:	_______________	Size:		________________

INSTALL SENSOR

Level Sensor ____ 	sensor feet ’locked’_______   (if you have questions, ask)

Declination:___________ Orientation:____________ (East Rod STS2; Brass Pin North CMG)

# Solar Power System Set-UP

# NOTE: The following tests should be performed with the solar panels in full sun, and with fully charged batteries.

1. Test output of the batteries (12.5 – 13 VDC WARNING: DO NOT test current). Voltage: ______

# 2. Connect the batteries to the power box

3. Test the voltage out of the power box to the DAS from pin A+ to C-  (Same as battery voltage.)

NOTE: Make sure the polarity is correct.  Voltage: _______

4. Test solar panel output (~18 Volts DC,).  Voltage: _______

5. Connect solar panels to power box

6. Test the voltage at the battery terminals (Greater than battery voltage above). Voltage: ____

INSTALL DAS

Connect GPS, and Sensor to DAS and then connect Power.

SENSOR Unlock

CMG-3T:     Use the host box to unlock the sensor. Press and hold both the Unlock and Enable Buttons for about 10 seconds.

CMG-3ESP: Remove the screw caps for each element and use the hex wrench to unlock all 3 elements. The red light on the host box will turn off when all of the elements have been unlocked.  Replace the screw caps before closing the sensor vault.

STS-2:	Use an STS-2 screwdriver to smoothly unlock all 3 elements. Give the STS-2 an initial centering pulse using the button on the host box.

DAS Setup

- Send Parameters to DAS (Edit Station Name & Enter Sensor serial number)

Work with Config ( load ( das_par_file ( Edit ( enter station name

(Channels(#1)(Details(enter sensor sn

(Send to DAS

(From DAS ( Edit ( Verify experiment name: _______________

2. Clear RAM (Control ( RAM ( Clear)

3. Reset System (Control ( Reset)

4. Format Flash Disk (Control ( Format Disk)   Disk 1:______   Disk 2:______

5. Sensor Check (Control ( Aux. Control ( Aux Ch.   Should be between +/- 1.5 Volts)

Check Mass Position Offsets

Ch 1:	___________V      Ch 2:   _________V      Ch 3:__________V

6. Monitor/Tap Test (Control ( Monitor ( Stream 1  [if >= 20 SPS] ( chans)

Ch 1:______		Ch 2: ______		Ch 3: ______

7. Check Clock Status (Control ( Status ( GPS) Time:	_________________________

Sec since LL:	_________________ *Note clock MUST lock before starting acquisition

Phase Diff:___________________ us (should be a small number)

SV’s: _____________    MODE > cycled

GPS Location of Site: ____________________________________________________

8. Start Acquisition (Control ( Status ( Start Acq)

Start time:	____________________________________(check year and time)

9. Verify RAM Increasing (Control ( Status ( Update)					Yes / No

10. Force RAM Dump to Disk 1 (Remove disk 2; Control ( RAM ( Dump)

Verify RAM decreases and disk 1 increases (Control ( Status ( Update)		Yes / No

Replace disk 2                                                         					Yes / No

11. Disk Setup  (Control (Disk)

Dump Threshold:  	66%

Auto-wrap 		No  (select the pull-down arrow to change setting)

Dump on ET		No

Tap the SEND button to send the information to the DAS

12.  Write .CFG File to Disk (Control(Status(DAS LP/WP)

Tap the WRITE button, then OK, to write the .cfg file to the disk.

Verify that the value of disk space used increases (Control (Status (Update).	Yes / No

13. DAS status ((Update) (or Control ( Status)

Acq: Start On / Off

Events:  ________			RAM:	 ________ of ________

Disk 1:  ________ of ________	Disk 2: ________ of ________

Temp:   ________			Power: Input ________ Bkup : ________

Ch:        ________			DS: ________

Firmware Version________ (Control(Satus(Versions)

14. Mass Position Offsets (recheck) (Control ( Aux. Control ( Aux Ch.)

Ch 1:	___________V      Ch 2:   _________V      Ch 3:__________V

Make sure all unused connectors are capped.

RECORD DIRECTIONS TO SITE AND ANY OTHER COMMENTS AND NOTES

TAKE A PICTURE OF THE SITE (Or Several)

CONTACT/LANDOWNER:______________________________________________________________

LOCATION OF SHIPPING CASES______________________________________________________

SITE NOTES: